## Slide 1: BL29 BOREAS Resonant X-ray Absorption and Scattering
BOREAS

## Slide 2: Soft XAS/XMCD & RMXS beamline
Scientific case
> Apple II undulator + VLS gratings mono  80-4000 eV energy range
> 2 endstations - HECTOR 6 T cryomagnet & MARES chamber+ 2 T magnet
Outline
1) BL performance
2) Endstations status
3) Users’ output

## Slide 3: BL performance - monochromator
1st und. harm.
3rd
5th and beyond
@ 100 mA
> April’14 : SM1 pitch blockage and SM double piping (DP) leak, friction (E reproducibility)
> March’13 : grating exchange motion friction (risk of leak opening)
Issues & Interventions (by Toyama)
> end 2014 (foreseen) : grating DP leak (Temp. stability)

## Slide 4: BL performance – KB mirrors
Intervention on april ‘14
> Blocked translations fixed (replacement of damaged IRELEC motors)
> Beam spot size not tunable before april’14  300 (H) x 800 (V) um (fixed)
Variable beam spot size from april’14. Best focusing  80 (H) x 30 (V) um
YAG crystal on STM plate (HECTOR sample holder)

## Slide 5: BL performance - E calibration & refs.
Ti L2,3
Niobium L23
Ruthenium L23
Ag L2,3
Calcium K-edge
Fe L2,3
Nb L2,3
S K
C K
Ru L2,3
Ca K
Photon energy
-
+

## Slide 6: BL performance - continuous scans
> High sampling frequency  larger density of pts/eV
> Less time between scans  better reproducibility
> Faster scans  more efficient use of beamtime
> Pending: implementation for “antiparallel mode” (i.e. linear intermediate pol.)
Variable step (open loop) und.+mono movement Total scan time  18 min.
Fe metal L2,3 XMCD
Fe metal L2,3 XMCD
Step scan
Continuous scan
Continuous und.+ mono movement.
Total scan time  2 min. (same resolution as step mode)

## Slide 7: BL performance - TOP UP tests
I0 FFT signal vs time (x103)
injection
> During injection we see noise increase @ 3.125 Hz due to kickers operation
> Limited impact on BL op. since number of affected scans is small
> The injection process ~10/20 s usually affects small portion of the spectra (typical scans take ~2 min)
Conclusion : No gating needed so far at BOREAS for TOP-UP operation
Freq. (Hz)

## Slide 8: BL endstation 1 - XMCD cryomagnet
1st endstation, “HECTOR” SC 6T/2T/2T magnet with VTI
> Running user experiments since May 2012 (29 till today, >5 days long in average)
> Performs according to specs (By up to 6 Tesla, 1.5 K<T <370 K)
> Detection modes: TEY + TFY, transmission
Recent upgrades
> Improved vacuum scheme (safe if blackout)
> Broadened remote macro-controlled operation
> Motorization of transmission diode arm
> Auto magnetic field-on flashing lights
In progress upgrades
> Fast loadlock (up to 4 sample holders simultaneously)
> Sample bias (to enhance TEY with X, Z- applied field)
> Anti-ice / condensation barriers
> Definitive Scaffold (for LN, LHe refill…)
> PFY detection (collaboration agreement with MPI Dresden under study)
> Dock for travelling suitcase, STM (coll. agr. with ICN2 under study)
Growing equipment @ preparation chamber
> Heating stage, evaporators, wobble stick, gas lines, ion gun, quartz balance, mini LEED
> New large turbo pump + ion pump  base pressure  10-10 mbar

## Slide 9: BL endstation 2 - Magnetic Scattering
Two detector arms
(IH-designed, manufacturing)
2 T magnet in/out tool (manufacturing)
Magnet rotation by McAllister (commissioning)
> Assembly, first tests with point detector
> Cabling, control (pool) in progress
> CCD commissioning in progress
> Final assembly and commissioning mid-late 2014
> Open to users from feb. 2015 (i.e. available in currently open call)
> Availability : reflectometry exps. without applied magnetic field
“MARES” endstation
Provisional support (definitive one late 2014)

## Slide 10: BL endstation 2 - detector arms
> Arm 1 – channeltron, microchannel plate, diodes, slits
> Arm 2 – XCAM CCD (SAT OK, june’14)

## Slide 11: BL endstation 2 – first spectra
[CoSi(5nm)/Si(3nm)] ML test sample reflectivity
> March 2014 IHR/Commissioning beamtime
Diode detection on temporary arm support

## Slide 12: Interstations transfer system – STM coupling
user STM-chamber
x-ray beam
Interstation sample transfer
XMCD
Vector-magnet
Scattering endstation
> In house-designed support (ready to assemble)
> System in service late 2013 / early 2014 (expected)

## Slide 13: Other developments - upgrades
> Small UHV prep. chamber, coupling to interstations transfer system
> “OESS” (off-endstation samples stage) for quick exchange of dirty samples
> Motorized system for (more) E reference samples
> Surface science instrumentation for MARES (mini-LEED head, linear translator…)
equipment
> Deeper macro / GUI control of vectorial magnetic field & sample temperature
> New macros for cooling down/warming up VTI, valves & EMs control…
> Long macros (overnight…) stability
> MARES macros
controls

## Slide 14: BL users operation 2013-2014
> Jan. 2013– Mar. 2014 : 14 user experiments performed (14 beamtime granted user proposals from 2012 call)
> May 2014 – Dec. 2014 : 13* accepted proposals (5 so far carried out) from 2013 call
* 7/13 require in-situ surface preparation (single crystal cleaning, evaporation etc.)
> Cold He-leak in VTI tail, drain current and 4-point contact wires, sample heater)
> Oct. – Nov. 2013 : liquid He supply shortage (global issue)  some experiments cancelled, beam profited to: 1) perform friendly users RT/zero B-field measurements, and 2) improve continuous scan mode
User experiments in HECTOR
Major issues / technical interventions in HECTOR

## Slide 15: IHR experiments – first publications
2012 beamtime

## Slide 16: User experiments – first publications
Fe
Co
2012 beamtime

## Slide 17: User experiments – PhD thesis contribs.
2012 beamtime
> High T SC with embedded nanoparticles
E. Solano et al, UAB-ICMAB-ALBA (PhD thesis, publication in prep.)
J. Padilla-Pantoja et al (in press, PhD thesis in prep.)
2013 beamtime

## Slide 18: IHR projects - collaborations
> Di-Fe (III) chiral complex molecular magnets (preliminary measurements)
Cl L1 edge ( 200 eV)
(20 times flux attenuation)
Fe L2,3
N K ( 400 eV)
C K ( 290 eV)
J. Giner-Planas, F. Di Salvo, M. Y. Tsang, S. M. Valvidares, P. Gargiani, J. Herrero-Martín, F. Teixidor, C. Viñas, N. Aliaga-Alcalde, E. Ruiz, D. Choquesillo-Lazarte,
ICMAB-CSIC, Univ. Barcelona, IACT-CSIC (Granada), ALBA

## Slide 19: Users evaluation (1/2)
Chart: BOREAS Beamline
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Quality of the beam | 1 | 0 | 0 | 4 | 12 | 6
Initial beamline status when you start your experiment | 1 | 0 | 3 | 4 | 13 | 2
Operation condition of the beamline | 1 | 1 | 3 | 5 | 10 | 3
User-friendliness of the beamline software | 1 | 2 | 3 | 4 | 11 | 2
Chart: BOREAS Beamline Staff
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Support received from beamline staff | 1 | 0 | 0 | 0 | 4 | 18
Proposal preparation | 5 | 0 | 0 | 0 | 15 | 3
Technical set-up preparation | 1 | 0 | 1 | 1 | 14 | 6
Beamline training | 1 | 0 | 0 | 5 | 9 | 8
Experiment run | 1 | 0 | 1 | 7 | 11 | 3

## Slide 20: Users evaluation (2/2)
Chart: 
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Access and support of present laboratories | 1 | 0 | 0 | 3 | 11 | 8
Chart: How does this beamline compare to others elsewhere in the world?
| No basis for comparison | Unfavorably | Favorably
BOREAS | 3 | 4 | 16
Users’ comments
> Bad pressure conditions in analysis chamber
> Some of the equipment is new and needs to be debugged
> J. Herrero, P. Gargiani and M. Valvidares were supporting us the whole time

## Slide 21: Acknowledgements
BL29 Staff
M. Valvidares, P. Gargiani (BL scientists),
J. Moldes (controls), A. Crisol (eng.), F. Farre (tech.),
X. Fariña and X. Serra (electronics)
Special ALBA contributors
E. Pellegrin, S. Ferrer, C. Colldelram, J. Nicolas, D. Fernandez, G. García, M. A. García-Aranda, L. Campos, User Office
Former staff & technical help
A. Barla, R. Martin, J.F. Moreno, D. Bacescu, J. C. Martinez, J. Lidon, Y. Nikitin, O. Matilla, S. Rubio, F. Becheri, J. Navarro, Ll. Ginés, A. Pascual, C. Ortiz, M. Llonch, C. Ruget, Alex, Paco, Eshraq, Xavi P. , C. Orozco, D. Iriarte, Thiago, A. Ruz, S. Astorga, Bern, Oscar, David, J. Ferrer, legal section (CFTs)
Externals: J. B. Kortright, N. Brookes, J. Criginski, C. Quiros, J. M. Alameda, C. Piamonteze, T. Schroder
External companies
TOYAMA, SCI-MAG, CVT, VG, HTS-110, PINK, SPECS, DODECON, FERROVAC, GAMMVAC, TECNOVAC, VAT, IBERLASER, HOSITRAD, VCS, AVS, AVACTEC,VAQTEC, MCALLISTER, HAMAMATSU, XCAM, AMBAR, INSERTY, SPARK,NANOTEC, SJUTS, KONIK, MDC, TRINOS, VAT