## Slide 1: Beamline specifications
1
Report on BL11-NCD
Sac 18, 1st July 2014

## Slide 2: Beamline specifications
2
Wavelength (Energy) range | 0.9 - 1.9 Å ( 6.5 -13 keV)
Flux at sample | >1.25 1011 ph/s 1.24 Å for a beam current of 250 mA
Bandpass (ΔE/E) | < 10-4
Beam size at sample | Variable between ~65 - 1200 μm horizontally ~30 - 265 μm vertically
Beam divergence at sample | <0.5x0.1 mrad2
Q range SAXS | 0.0066-0.7 Å-1
2Ɵ range WAXS | 3.0° - 62°

## Slide 3: Optics Hutch - Monochromator
3
Chart: Theoretical Bragg Angle (°)
| 13.699840380843847 | 13.686471977575392 | 13.728336003359132 | 13.708209035777527 | 13.689811607400733 | 12.720084759495409 | 12.701966550013369 | 12.6892001057759 | 12.657468225202422 | 12.723397679714468 | 12.707569002110807 | 12.691779994935208 | 11.801837858251639 | 11.790705136198115 | 11.767272060980417 | 11.814232413286394 | 11.80555347313868 | 11.793177245671442 | 11.775894375534047
 | 13.7151 | 13.7014 | 13.7441 | 13.7225 | 13.706 | 12.7339 | 12.7155 | 12.7012 | 12.6687 | 12.7352 | 12.7185 | 12.7037 | 11.8146 | 11.8036 | 11.7823 | 11.8267 | 11.8175 | 11.807 | 11.7893
Slope: 1.0010
Intercept: 0.0006
Bragg Angle calibration

## Slide 4: Optics Hutch - Monochromator
4
Chart: 1/cos Ɵ
| 1.0256465727166462 | 1.0250583320948439 | 1.0244903390900464 | 1.0239416645357462 | 1.023411432526381 | 1.0228988167589292 | 1.0224030371670543 | 1.0219233568211361 | 1.0214590790702553 | 1.0210095449046088 | 1.0205741305189784 | 1.0201522450597944 | 1.0197433285400228 | 1.0193468499076301 | 1.0189623052547392 | 1.0185892161557921 | 1.0182271281241355 | 1.0178756091774057 | 1.017534248502977 | 1.0172026552155213 | 1.0168804571994279 | 1.0165673000294853 | 1.0162628459637915 | 1.0159667730033852 | 1.0156787740135627 | 1.0153985559022645 | 1.0151258388513149 | 1.0148603555966342 | 1.0146018507538697 | 1.0143500801861827 | 1.0141048104111856 | 1.0138658180442657 | 1.013632889275756 | 1.0134058193796023 | 1.0131844122513698
 | -1582.3125 | -1583.6175 | -1585.575 | -1589.49 | -1589.49 | -1591.4475 | -1593.405 | -1594.71 | -1596.015 | -1597.32 | -1598.625 | -1597.9725 | -1601.235 | -1602.54 | -1602.54 | -1604.4975 | -1604.4975 | -1606.455 | -1606.455 | -1608.4125 | -1608.4125 | -1610.37 | -1610.37 | -1612.3275 | -1613.6325 | -1612.98 | -1616.2425 | -1615.59 | -1616.895 | -1616.2425 | -1619.505 | -1620.81 | -1620.1575 | -1621.4625 | -1621.4625
 | -1622.7675 | -1626.6825 | -1626.6825 | -1628.64 | -1630.5975 | -1629.945 | -1629.945 | -1631.9025 | -1635.165 | -1634.5125 | -1633.86 | -1635.165 | -1634.5125 | -1635.8175
2nd Crystal pitch Calibration

## Slide 5: Undulator
5
ID gap = 5.8mm
Experimental data
Theoretical data calculated with no offset in the front end aperture.
Energy scan

## Slide 6: Undulator
6
Chart: Energy (keV)
| 8.9 | 9 | 9.1 | 9.1999999999999993 | 9.3000000000000007 | 9.4 | 9.5 | 9.6 | 9.6999999999999993 | 9.8000000000000007 | 9.9 | 10 | 10.1 | 10.199999999999999 | 10.3 | 10.4 | 10.5 | 10.6 | 10.7 | 10.8 | 10.9 | 11 | 11.1 | 11.2 | 11.3 | 11.4 | 11.5 | 11.6 | 11.7 | 11.8 | 11.9 | 12 | 12.1 | 12.2 | 12.3
 | 5.92 | 6.0001 | 6.0702 | 6.1303 | 6.2036 | 6.2736 | 6.3436 | 6.4138 | 6.4837 | 6.5537 | 6.6238 | 6.6937 | 6.7637 | 6.8338 | 6.9037 | 6.9737 | 7.0639 | 7.1238 | 7.1937 | 7.2637 | 7.338 | 7.4079 | 7.4793 | 7.5492 | 7.6193 | 7.7002 | 7.7702 | 7.8403 | 7.9102 | 8.0005 | 8.0702 | 8.1303 | 8.2204 | 8.2803 | 8.3704
 | 5.9 | 5.9502 | 6.0003 | 6.0634 | 6.1004 | 6.1502 | 6.2002 | 6.2641 | 6.3044 | 6.3446 | 6.4047 | 6.4449 | 6.5049 | 6.5648
Undulator gap calibration

## Slide 7: Energy pseudo-motor
7
Done:
Calibration of the Bragg angle of the DCM
Calibration of the 2nd crystal pitch
Calibration of the undulator gap for the 5th and 7th harmonic
To be done :
Calibration of the undulator gap for the 3rd harmonic
Calibration of the vertical translation of the 2nd crystal

## Slide 8: Upgrade – Air conditioning
8
February 2014

## Slide 9: Upgrade – Air conditioning
9
June 2014 (After fans installations)
Proper air conditioning to be installed by the end of this year

## Slide 10: Detector - Rayonix
Waxs detector
10
Active image area: 255 x 85 mm2
Total number of pixels: 5760 x 1920
Pixel size ( not binned): 44 µm
Frame rate: 10 – 140 frames / sec
Global count rate: unlimited
Dynamic range: 16 bits

## Slide 11: Detector - ImXPAD
Saxs detector
11
Active image area: 152 x 149.6 mm2(H x V)
Total number of pixels: 1200 x 1120 pixels2 (H x V)
Pixel size ( not binned): 130 µm
Frame rate: 250 frames / sec
Global count rate: >1011 photons/sec
Site Acceptance Test planned for September 2014
Factory Acceptance Test done in May 2014

## Slide 12: Upgrade – Window protection
12
Connected to the photon shutter
Possibility to open it manually if the vacuum is broken in the pipe
Light on doors => Safe to enter in the EH
Available in September 2014

## Slide 13: Upgrade – WAXS nose cone
13
Available in April 2014

## Slide 14: Upgrade – On-Axis Camera
14
1.67X \ 120FL Lower lens 35-00-04-000 | WD 155mm, Min Zoom: Mag - 0.47X; Res-22.2m; DOF-2.5mm; 10mm x 7.7mm Max Zoom: Mag - 5.80X; Res-4.2m; DOF-0.089mm; 0.83mm x 0.63mm
1mm
0.75mm
10mm
7.25mm
Available in September 2014

## Slide 15: Upgrade – On-Axis Camera
15
Available in September 2014
Different lighting system depending on the contrast of the sample

## Slide 16: Upgrade – Schedule
16
Project | Status | Delivery
Flight tube autom. protection | Under production | 20/06/2014
WAXS + Nose cone sup. Stages | Conceptual Design | 22/04/2015
Sample platform |  | 29/05/2015
SAXS detector table |  | 15/04/2015
Microfocus lenses |  | 01/04/2015
BCE fast shutter/camera |  | 27/05/2015
DCM foils |  | 08/07/2015
Beam stopper |  | 01/07/2015
KF support pipes |  | 30/06/2015
Flight tube cable chains |  | 14/10/2015
Support various |  | 28/10/2015
KB system |  | 30/08/2016

## Slide 17: Control – New projects
17

## Slide 18: Control – pyFAI implementation
18
SampleDetDist= 1.260982e-01m
PONI= 4.351847e-02, 1.918830e-01m
rot1=-0.556464 rot2= -0.000100 rot3= -0.000000 rad
DirectBeamDist= 148.503mm
Center: x=6143.644, y=988.719 pix
Tilt=31.883 deg
tiltPlanRotation= -0.011 deg
J. Kieffer, PyFAI, Journal of Physics: Conference Series 425 (2013)
Source code available from ESRF

## Slide 19: Control –New features
19
Save environment information in image headers
  Basic image information: Exposure time, time stamp and readable image date time. (End of March 2014)
  To have this information in the ASCII file(End of March 2014)
  First energy, machine current, ion chamber current (End of March 2014)
  Next will be to have other values as temperatures for linkam scans
  Final solution will be a way to set a list of attributes to be saved in image headers (End of summer shutdown 2014)
  To have Rayonix information like time stamp in the header

## Slide 20: Users - 2013
20
8 users groups
7 publications
but 1 group is publishing 3 articles after 5 days of beamtime

## Slide 21: Users feedback 2013
21
Chart: NCD Beamline
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Quality of the beam | 1 | 0 | 0 | 0 | 1 | 12
Initial beamline status when you start your experiment | 2 | 0 | 1 | 0 | 2 | 9
Operation condition of the beamline | 2 | 0 | 1 | 1 | 4 | 6
User-friendliness of the beamline software | 2 | 0 | 0 | 3 | 7 | 2
Chart: NCD Beamline Staff
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Support received from beamline staff | 2 | 0 | 0 | 0 | 2 | 10
Proposal preparation | 6 | 0 | 0 | 0 | 5 | 3
Technical set-up preparation | 3 | 0 | 0 | 0 | 6 | 5
Beamline training | 3 | 0 | 0 | 0 | 3 | 8
Experiment run | 2 | 0 | 0 | 0 | 5 | 7

## Slide 22: Users feedback 2013
22
Chart: How does this beamline compare to others elsewhere in the world?
| No basis for comparison | Unfavorably | Favorably
NCD | 7 | 0 | 7
Chart: 
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Access and support of present laboratories | 3 | 0 | 0 | 1 | 2 | 8

## Slide 23
23
The role of decorin in organizing the extracellular matrix
Christina Kamma-Lorger (PI), Marc Malfois, Juan Carlos Martinez, Eva Crosas, Keith M Meek, Carlo Knupp, Cecilie Bredrup, Eyvind Rødahl, Alex Peralvarez

## Slide 24: In house project 2014 An investigation of structural changes of fibrillar skin proteins and microRNAs in cancer Christina Kamma-Lorger (PI) and Marc Malfois
24
Skin Cancer
Study skin collagen and keratin ultrastructure in squamous cell carcinoma using SAXS and EM 3D tomography
Human miRNA structure using BioSAXS solution scattering
Current progress:
Internal funding was granted for 2014. Installation of cell culture laboratory facility.

## Slide 25: Study of radiation damages on Macromolecules – A NCD/XALOC project E. Crosas, C. Kamma-Lorger, J.C. Martinez, M. Malfois, A. Castellvi, F. Gil, J. Juanhuix
25
Funded by the Spanish Safety Nuclear Council (CSN)
Radiation damage is observed as aggregation of the protein in the small angle S region, affecting size and shape determination of the protein.
Figure
Effect of radiation damage on lysozyme. Open circles: short-exposure (165 Gy); solid circles: long-exposure (3300 Gy). The upward curvature in the small S region (solid circles) points to the presence of radiation-induced aggregates.
From Kuwamoto S., J Synchrotron Radiat. 2004 Nov 1;11(Pt 6):462-8

## Slide 26: Study of radiation damages on Macromolecules – A NCD/XALOC project E. Crosas, C. Kamma-Lorger, J.C. Martinez, M. Malfois, A. Castellvi, F. Gil, J. Juanhuix
26
Systematic comparison of radiation damage on several proteins using both SAXS and MX techniques
Test different pH conditions and scavengers, including new potential compounds against radiation damage
Correlation of SAXS and MX results with functional studies such as:
  Mass Spectrometry MALDI-TOF
  Enzyme activity assay
Beamtime awarded:
ALBA
ESRF and Diamond to take advantage of the BioSAXS robot.
For more information, see the related poster tomorrow

## Slide 27: Liposomes in gene therapy: Lipoplex structure-biological activity relationship A. L. Barrán-Berdón, B. Yélamos, M. Malfois, E. Aicart and E. Junquera
27
Submitted in Langmuir
Lipofection is a technique used to introduce genetic information (DNA or siRNA) into a cell by means of lipid nanostructures that can easily fuse with the cell membrane.
Lipoplex (lipids + DNA or siRNA) structure is known to clearly affect transfection efficiency (TE).
+
Transfection Efficiency
(100)
(200)
(300)
(10)
(11)
(20)
(21)
(30)
(22)
(31)
-

## Slide 28: Liposomes in gene therapy: Lipoplex structure-biological activity relationship A. L. Barrán-Berdón, B. Yélamos, M. Malfois, E. Aicart and E. Junquera
28
Submitted in Langmuir
Lipoplex structure depends on the molecular characteristics of the lipids used.
Lipoplexes showing inverted hexagonal phases (HII) are more efficient on transfection than lipoplexes with a lamellar ones (L). Hexagonal phases are more fusogenic; accordingly, after endocytosis, they release more easily DNA, from the endosome to the cytosol.
SAXS together with transfection assays reveal as a potent tool to correlate lipoplex structure with transfection levels (structure-biological activity relationship).
GFP expression

## Slide 29: Application of SAXS to the structural analysis of membrane vesicles Oscar Moran, Istituto di Biofisica, CNR, Genova, Italy
29
Published in Eur. Biophys. J.
Submitted in Cellular and Molecular Life Sciences
From scattering theory, analytic expressions are derived for the bilayer form factor over a spherical geometry, assuming the lipid bilayer electron density to be composed of a series of Gaussian shells.
Microsome membranes unilamelar vesicles
Electron density profiles of from microsome vesicles wall of native cells (green) and cells overexpressing CFTR (blue; protein involved in cystic fibrosis). Micrographs of vesicles are shown.
Sketch of the model of a vesicle consistent with the measured SAXS data.
SAXS spectra of the WT-CFTR membranes
The electronic density profile was calculated from the five Gaussian model
The decomposition of each singular Gaussian used to model the electronic density

## Slide 30: Application of SAXS to the structural analysis of membrane vesicles Oscar Moran, Istituto di Biofisica, CNR, Genova, Italy
30
Large unilamellar liposomes (LUV) obtained by extrusion of phospholipids, treated with drugs for the cystic fibrosis therapy.
untreated
+ VX-770
+ VX-809
Electron density of the LUVs wall showing the destabilization induced by drugs.
Partially supported by the Italian Cystic Fibriosis Foundation
Published in Eur. Biophys. J.
Submitted in Cellular and Molecular Life Sciences

## Slide 31: Carbon nanotube - 1-
Controlling carbon nanotube type in macroscopic fibers synthesized by the direct spinning process Reguero et al. accepted at Chemistry of Materials
Synchrotron XRD confirms the increase of graphitic layers at turbostratic separation
S ↑ in feedstock  Intensity ratio [(100)Graphite / (111)Fe]
The research group has a process to produce kilometres of continous macroscopic fibres of CNTs
Low S  SWCNTs
High S  MWCNTs
They can tailor thee type of nanotubes through the addition of sulphur precursor (thipohene)

## Slide 32: Carbon nanotube - 2-
The CNT fibres have a complex hierarchical structure.
Its mesoporosity makes liquids infiltrate the fibre,
which changes its eletrical conductivity.
This non-Ohmic behaviour is a manifestation of
nanoscale effects observed on a macroscopic scale.
It arises due to fibre swelling, which can be observed by SAXS.
Electric Field-Modulated non-Ohmic Behavior of Carbon Nanotube Fibers in Polar Liquids (Terrones et al., ACS Nano, submitted)

## Slide 33: Beamline staff
33
Beamline Responsible:
  Marc Malfois
Beamline scientist:
  Christina Kamma-Lorger
  Juan-Carlos Martinez
Postdoctoral Research Associate:
  Eva Crosas
Technician:
  Alejandro Enrique (Left June 2014)
Controls:
  Gabriel Jover
Engineering:
  Joaquin Gonzalez
  Carles Colldelram
Electronics:
  Abel Fontsere