## Slide 1: Machine Status and Top Up Operation
Francis Perez

## Slide 2: Outline
Operation
  2013
  2014 (so far)
Top Up
Accelerator Developments
  Injection bump closure
  FOFB
  Beam lifetime
  Energy measurement
  Interferometry beam size
  “67 MeV” injector
2

## Slide 3: Operation
ALBA Accelerators
Operation 2013
Operation 2014 (so far)

## Slide 4
ALBA Accelerators
LINAC
100 MeV
BOOSTER
100 MeV – 3 GeV
STORAGE RING
3 GeV
ALBA Accelerators

## Slide 5
SR and BOOSTER sharing the tunnel
Booster
SR
In vacuum undulator
RF cavities
Bending
ALBA Accelerators

## Slide 6
Storage Ring: Main Parameters
Electron beam energy 3.0 GeV
Storage Ring Circumference 268.8 m
Number of cells 16
Symmetry 4
Straight sections: 4 Long 8.0 m for 3 ID & Injection
12 Medium 4.4 m for 12 ID
8 Short 2.6 m for 2 ID & RF & Diagnostics
Beam current 250 mA
Emittance 4.3 nm.rad
Lifetime > 10 h
ALBA Storage Ring

## Slide 7
ALBA Storage Ring
E [GeV] | 3
C [m] | 268.8
Qx | 18.19
Qy | 8.24
x [nm rad] | 4.3
ξx | -34
ξy | -24.5
αp | 8.4×10-4
a2 | 2.5×10-3
E/E | 1.05×10-3
Jx/Jy/JE | 1.3/1/1.7
tx/ty/tE [ms] | 4.0/5.3/3.1

## Slide 8
The cooling incident force a major modification of the operations calendar in July 2013
Total hours lost (runs 03 to 05): 1416 h
Total hours added (runs 06 to 09): 824 h
Operation 2013

## Slide 9
The cooling incident force a major modification of the operations calendar in July 2013
 | Initial schedule | Revised schedule | Delivered
BL(2012) |  |  | 2387.3
BL(2013) | 3539.5 | 3069.0 | 2971.5
Injections | 252.5 | 219.0 | 
M+csn | 1488.0 | 1368.0 | 
W | 2432.0 | 3240.0 | 
Off | 1048.0 | 864.0 | 
Operation 2013
Total hours lost (runs 03 to 05): 1416 h
Total hours added (runs 06 to 09): 824 h

## Slide 10
2 shifts/day
Beam availability
> 98%
> 97%
> 97%
> 91%
> 95%
> 97%
83.8 % average over initial schedule
96.8 % average over revised schedule
Operation 2013

## Slide 11
BL Beamlines PSS
CT Controls
DI Diagnostics
EPS EPS interlock
ID Insertion Devices
INF Infrastructure
LI Linac
MML High Level Software (MML)
OE Human error
OI Orbit Interlock
OT Others
PM Pulsed magnets
PS Power Supplies
PSS PSS (except BLs)
RF Radio Frequency
VC Vacuum
WI Water interlocks
Analysis of failures per sub-system
Total of 1109.5 h of beam downtime
Operation 2013

## Slide 12
BL Beamlines PSS
CT Controls
DI Diagnostics
EPS EPS interlock
ID Insertion Devices
INF Infrastructure
LI Linac
MML High Level Software (MML)
OE Human error
OI Orbit Interlock
OT Others
PM Pulsed magnets
PS Power Supplies
PSS PSS (except BLs)
RF Radio Frequency
VC Vacuum
WI Water interlocks
Analysis of failures per sub-system w/o cooling & MPW chamber replacement
Total of 97 h of beam downtime ≈ 0.7 h/BL day
Operation 2013

## Slide 13
BL Beamlines PSS
CT Controls
DI Diagnostics
EPS EPS interlock
ID Insertion Devices
INF Infrastructure
LI Linac
MML High Level Software (MML)
OE Human error
OI Orbit Interlock
OT Others
PM Pulsed magnets
PS Power Supplies
PSS PSS (except BLs)
RF Radio Frequency
VC Vacuum
WI Water interlocks
Analysis of failures per sub-system w/o cooling & MPW chamber replacement
Total of 123 failures causing a beam dump ≈ 0.9 failure/BL day  25 h MTBF
Operation 2013

## Slide 14
 | 2012 | 2013
Beam for BLs | 2387.3 | 2971.5
Availability | 85 % | 96 % *
MTBF | 21.0 | 25.0 *
MTTB(back) | 1.0 | 0.8 *
Summary
Operation 2013
* Excluding water cooling problem

## Slide 15
Operation 2014 (so far)
Beam availability:
RUN_01 96.7 %
RUN_02 97.1 %
RUN_03 97.6 %
RUN_04 97,7 %
41,2 h
11.0 h Water leak at FE34
2.3 h PSS intlk at BL09
1,1 h

## Slide 16
Operation 2014 (so far)

## Slide 17
FE absorber temperature – ok
Pinhole aluminium window water leak
Damaged
Kicker
Electronics

## Slide 18
Repaired the following day
Pinhole aluminium window water leak
?
Small (350um) hole in the Cu cooling pipe
Microscope image

## Slide 19
LINAC Gun High Voltage PS
Before starting run #3, during the subsystem tests,
the LINAC Gun High Voltage PS (-90 kV) did broke.
We did not have available spare:
Exchanged by -50 kV PS (from RF group) and commissioned the Linac with reduced e-gun beam energy.
BESSY lend us a -90 kV as back-up
(many thanks to Andreas that pick up the phone on May 1st and SOLEIL that offered also their spare)
Reparation already done and new spare purchased
Back to normal along this week

## Slide 20
IOTs – RF transmitter upgrade
IOTs Upgrade: 1 prototype end of 2014
4 series units summer 2015
In addition, more Thales IOTs spares purchased.

## Slide 21: Top Up

## Slide 22
Top-up operation test with users: last week
June 25th – 28th, 2014
Top Up

## Slide 23
Top Up
Top-up operation test with users: last week
June 25th – 28th, 2014

## Slide 24
Top Up
Top-up operation test with users: last week
June 25th – 28th, 2014
XALOC: xBPM(y)
Decay mode
Machine day
Top-up
(jumps are due to changes in the ID gap)

## Slide 25
    Influence of injection on BLs,
    after several test together with the BL scientists.
 | sensitive to BO | sensitive to KI
BL04 | No | 
BL09 | No | 
BL11 | No | 
BL13 | No | 
BL22 | No | 
BL24 | Slight increase of noise, but the effect is marginal | Yes, gating is needed
BL29 | No | Slight increase of noise but the effect is marginal. No gating is needed
Top Up
Soft gating provided to BLs

## Slide 26
Top Up
12.5 kHz
(close to CIRCE)
Acoustic noise -due to the Booster PS ramping at 3 Hz- at Experimental Hall
Technical solutions to damp the noise are under investigation
© Safety

## Slide 27
Radiation measured during Top-Up
Well below Public Zone & Similar to decay mode
© Safety
June 25th

## Slide 28
Radiation measured during Top-Up
These are beam loses, not Top Up related radiation, but triggered the Top Up PSS logic to act in order to maintain the dose below 2 uSv / 4h – Public Zone.
© Safety
June 29th
Top Up – FE open
M day – FE closed

## Slide 29
    Specific measurements at BOREAS Optical Hutch, only BL where higher doses were detected
From last SAC
Radiation measured during Top-Up
© Safety
Data taken for the
worst (not allowed) conditions

## Slide 30
Radiation measured during Top-Up
Only gammas detected
at BOREAS
Under investigation
© Safety

## Slide 31
Additional Shielding
BOREAS - BL29
© Safety

## Slide 32: Accelerator Developments
Injection bump closure
FOFB
Beam lifetime
Energy measurement
Interferometry beam size
“67 MeV” injector

## Slide 33
Top Up – Closing injection bump
  The BPM turn by turn data response matrix to the kicker voltage is measured at turn #9 (end of kicker pulse)
dX [mm] / Kicker dV [V]
Injection Bump Closure
  TbT orbit oscillations are corrected applying an SVD
  Nowadays, manual readjustment are done

## Slide 34
Top Up – Closing injection bump
Best achieved
Injection Bump Closure
Distortion
at
Medium
Straigth
Section
X
Y

## Slide 35
Fast Orbit FeedBack (Vertical)
FOFB OFF
FOFB ON
3 Hz
Notes: 35

## Slide 36
Fast Orbit FeedBack (Horizontal)
3 Hz horizontal
kicker induced noise
is amplified by FOFB
Notes: 36

## Slide 37
Fast Orbit FeedBack
FOFB is working but still:
  Fine tune required
  Robustness (long runs) to be tested
  Error handing
  Friendly user interface
Notes: 37

## Slide 38
Theoretical Lifetime Calculation
Project under development,
but we already take into account:
Current
Pressure
SCRV
RF Voltage
Energy Loss per Turn
Next step:
Improve Touschek Lifetime Calculation using Piwinski Model (Coupling, RF Voltage)
Lifetime calculator
Measured data
or
User defined values
Lifetime

## Slide 39
Good agreement with the measured Lifetime
Err < 3%
Lifetime
It is still a work in progress, some constants are manually adjusted

## Slide 40
Energy Measurement
Spin depolarization
2.97813 ±0.00002 GeV
3 different days
3 different RF frequencies
3 different energies: ±0.0001 GeV

## Slide 41
Interferometry beam size

## Slide 42
Vibration of the mirrors? Air turbulence?
Interferometry beam size
But…

## Slide 43
“67 MeV” injector
Objective:
Operate the LINAC with only one Klystron:
  Either KA1 or KA2, for redundancy
  Compatible with installation and test of the second one

## Slide 44
“67 MeV” injector
1st turn in Booster

## Slide 45
  SUMMARY
Accelerators operate reliable for users
    96% 2013 (excluding cooling problem)
    97% 2014 (so far)
Top Up operational, but
    Acoustic shielding of Booster PS
    Radiation local shielding at BL29
Objectives until end of 2014
    FOFB operational
    Injector “67 MeV”
    Increase operation current
Summary

## Slide 46
Thanks to you,
to the accelerators team
and many others
Division Head
F. PEREZ
Operations
M. PONT
Operators
F.FERNANDEZ
M. ALVAREZ
R. NUÑEZ
D. YEPEZ
P. SOLANS
RF & Linac
RF
A. SALOM
B. BRAVO
J. OCAMPO
Linac
R. MUÑOZ
D. LANAIA
IDs
J. CAMPMANY
Magn Lab
J. MARCOS
V. MASSANA
Beam Physics
U. IRISO
Beam Dynamics
G. BENEDETTI
T. GÜNZEL
Z. MARTÍ
M. CARLÀ
oPAC
Diagnostics
A. OLMOS
A. NOSYCH
L. TORINO
oPAC
J.C GIRALDO
M. SOS
N. AYALA