## Slide 1: Report on BL24- CIRCE
Lucia Aballe
Michael Foerster
Virginia Perez
Carlos Escudero
Support:
Jose Prieto (Technician, with CLAESS)
Nahikari Gonzalez (Mechanical engineer, with XALOC)
Abel Fontsere, Toni Camps (Electronics engineer, with NCD)
Fulvio Becheri (Controls, with ID)
Josep Nicolas (Optics, transversal section)
Eric Pellegrin (Section Head)
1

## Slide 2: BL24- CIRCE: Photoemission Spectroscopy and Microscopy
Status and progress: technical aspects
  Beamline (soft X-ray PGM with helical undulator)
  Photoemission electron microscope (PEEM)
  Near ambient pressure photoemission (NAPP)
Productivity: User Statistics, Output
2

## Slide 3
Beamline
3

## Slide 4: KB motor exchange (1/2014)
“Problems with KB optics need solution! Risk of loosing the station for ca. 8 weeks.”
Three linear actuators for HFM x, HFM pitch and VFM pitch were replaced during Christmas’13 shutdown.
Irelec provided and installed a new-designed actuator, UHV compatible and capable of standing larger loads (up to 200 N).
All motions are functional, friction effects still present but largely reduced.
Intervention:
Nahikari González
José Prieto / Francesc Ferrer
Marta Llonch / Jon Ladrera
Josep Nicolas
Oscar Borrego
Lluís Gines
Resolution test obtained during intervention
4

## Slide 5: HEG cleaning (in progress)
- Ex-situ characterization by AFM
- HEG ex-situ cleanning with oxygen plasma: in progress
- Post-cleaning characterization and reinstalling in PGM
HEG flux= MEG flux/4: Carbon contamination?
darkest area
cleanest area
5
Notes: Au coated Si, no binding layer
Cleanning tests in Au coated Si mirror samples.

## Slide 6: Top up tests – mirror drain current @ 10Hz
6

## Slide 7: Top up tests – beam spot on sample
7
Spot on sample (PEEM): no major change on size or position, but strong effect on intensity
10% dip, 3Hz, no difference with/out FOFB
vertical beam profile at sample

## Slide 8: Top up tests – spectra in PEEM
Normalization does not work exactly
So far, no I0 signal synchronized with NAPP analyzer software
Not yet solved: acoustic noise, mainly affecting PEEM
8

## Slide 9: Top up tests – XAS spectra in PEEM
9
Normalization seems to depend on timing/synchronization  faster aquistion needed

## Slide 10: Linear polarization using AP (Controls)
10
For information/expert user
Cross-implemented from Boreas (Fulvio Becheri)
Change parallel/antiparallel mode via macro (sets phase=antiphase=0, changes ID correction coils settings and pseudomotors (equations))
Future: single pseudomotors set (and panel)
Changed by user

## Slide 11
PEEM
11

## Slide 12: PEEM advances
Microscope:
  (Preliminary) Alignement and operation at 15kV
  Vacuum: Installation of Ti sublimators
Sample environment:
  (moderate) magnetic field cooling
  Electrical poling
Controls:
  XAS and XMCD previews available (much faster operation)
  Continous Camera readout in XMCD macro (saving 25-50% time, used in regular beamtime)
12
1 um
XMCD image taken at 15 kV (showing film grain structure)

## Slide 13: PEEM magnetic sample holders
  New sample holders with electro- magnets in use (IP and OOP fields) (engineering: N. Gonzalez, J. Ferrer, intern student Pol Alonso)
  Added polarity switch, debug current source and sample mount
13
XMCD images (FOV 50 um) showing in-situ switching (opposite remanent states) and image taken at 9mT
Based on design from BESSY (F. Kronast)

## Slide 14: Imaging eletric field induced magnetic switching at RT
14
D. Pesquera, B. Casals, G. Herranz, J. Fontcuberta (Institut de Ciencia de Materials ICMAB-CSIC)
La2/3Sr1/3MnO3//PMN-PT (100)
with OOP electrical poling (A. Fontsere & Electronics)

## Slide 15: Spectromicroscopy of the Fe3O4 (100) surface
15
L. Martin, M. Monti, J. Marco, J. Figuera (IQFR-CSIC, Instituto de Química Física "Rocasolano")
Top: LEEM, DF-LEEM, XMCD-PEEM, XMLD-PEEM (FOV 20 um)
Bottom: XMCD and XMLD spectra extracted from opposite domains

## Slide 16: Reactive growth of cobalt germanide islands on Ge (001)
16
T. Schmidt, I. Heidmann, M. Ewert, J.I. Flege, T. Grzela, T. Schroeder and J. Falta (Institute of Solid State Physics, University of Bremen, & IHP Microelectronics GmbH, Frankfurt (Oder), Germany)
µ-LEED on one needle (27 V)

## Slide 17: Magnetic depth profiling of nanowires
17
S. da Col et al., Phys. Rev B 89, 180405(R) (2014)
Co L3
Co L2
Co L3
„preedge“
1 um
PEEM direct images with photon energy around Co L3
C. Bran, A. Asenjo, R. Perez, E. Palmero, M. Vazquez (ICMM-CSIC) and A. Fraile (UB)

## Slide 18: For the future: using in-plane electrodes
18
With I. Fina (MPI Halle, ICMAB-CSIC)

## Slide 19
NAPP
19

## Slide 20: Peltier cooling manipulator delivered
Sample T(K)
Peltier Current (A)
Time (s)
1000
2000
3000
4000
0
245
250
255
260
265
270
275
280
285
290
T (K)
0.25
0.50
0.75
1.00
1.25
1.50
1.75
I (A)
FAT Dec 2013
In UHV
SAT April 2014
Lowest T in UHV: -22.3 oC in 10 min
Lowest T at P=1 mbar : -13 oC in 10 min
20

## Slide 21: Analyzer commissioning
- NAPP Prelens very sensitive to low retarding ratios
Remarkable count rate increase:
issue with slits solved/nozzle exchanged
21

## Slide 22: C&C and Electronics Highlights
Heating remote control with button heater. With laser heater, in
progress (A. Fontsere)
Implementing NEXAFS scans with NAPP analyzer in Auger mode
(F. Becheri and R. Homs)
Safety vacuum interlock system and vacuum tank system
(Electronics, Engineering)

## Slide 23
23
Puregas system
  Work under dynamic conditions with up to three different gases simultaneously (flow control from 0 to 30 mL/min for each gas)
  Gas lines for specific gases:
1) H2 and hydrocarbons
2) inert gases
3) O2/CO2
Electro-polished pipe with the possibility to insert additional purifiers: low contamination level, down to ppb’s
In comissioning

## Slide 24: CEM system (Controlled Evaporator and Mixer)
24
In comissioning
(Support by Francisco Trujillo)
  It enables to mix liquid flows of 0,25-5 g/h resulting in saturated vapour flows of up to 60 ml/min.
This system replace bubblers allowing more control

## Slide 25: Liquid microjet system
25
  Segregation phenomena of ions in solutions
  Adsorption of gas molecules on liquid films
  Electronic structure of biomolecules in water and other organic molecules in different solvents minimizing the potential beam damage
  Synergy with project for XFEL
Ordered

## Slide 26: RhPd-based catalysts for hydrogen production (UPC Barcelona)
C2H5OH + 3 H2O è 6 H2 + 2 CO2
Ethanol steam reforming
for H2 generation
Pd and Rh 3d levels during reaction at different probe depth
N. J. Divins, I. Angurell, C. Escudero, V. Perez-Dieste and J. Llorca (Submitted)
Rearrangement of metals in NP´s
during catalytic reaction?
26

## Slide 27
Breathing chemistry
By courtesy of Dr. Celia Rogero (CSIC-UPV/EHU)
7 x 10–10 mbar
1 x 10–2 mbar (O2)
1 x 10–9 mbar
27

## Slide 28: Next steps/goals:
28
BL: Beam chopper and normalization at Carbon edge
BL: Continous energy scan (adapt from BOREAS)
NAPP: Commissioning of CEM and liquid microjet system
NAPP: Installation of a conventional X-ray source
NAPP: Integration of I0 signal with SPECS software (NEXAFS & Top-up)
PEEM: reduce camera deadtime for single image (XAS/XPS)
PEEM: in-plane electrodes and current injection
PEEM: HF cable connections and electronics
Maintain oversupscription & develop user communities (e.g. workshops)
Solve contractual situation of BL staff (2 of 4)
BL- Scientists: Find more time for „scientific“ work (demanding experiments)
BL- Users: output in form of papers

## Slide 29
29
Productivity and Output

## Slide 30: User statistics – Call 2014
(from 2014 Call Statistics provided by User Office)
30
  Total 47 Proposals submitted: 810 shifts
  We foresee to be able to give 225 shifts to users only during the call period (5-12/2014) (NAPP and PEEM together)
  High oversubscription can cause frustation/pressure

## Slide 31: User statistics: 2014 and before
User feedback
 | PEEM – 2014 (May-December) | NAPP – 2014 (May-December) | PEEM – since 11/2012 | NAPP – since 9/2013
Beamtime (d) | 38 | 37 | 51 | 12
Proposals (Groups) | 8(7) | 7(7) | 8(4) | 2(2)
Friendly (d) | 0 | 0 | 30 | 18
Follow-up proposals | 3 | 3 | - | -
31
  Receiving LEEM users and in-House collaborators
Countless days of experiment preparation & in-situ sample preparation

## Slide 32
User feedback
32
Chart: CIRCE Beamline
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Quality of the beam | 0 | 0 | 0 | 2 | 3 | 3
Initial beamline status when you start your experiment | 0 | 0 | 0 | 2 | 4 | 2
Operation condition of the beamline | 0 | 0 | 0 | 2 | 4 | 2
User-friendliness of the beamline software | 0 | 0 | 1 | 2 | 4 | 1
Chart: CIRCE Beamline Staff
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Support received from beamline staff | 0 | 0 | 0 | 0 | 0 | 8
Proposal preparation | 0 | 0 | 0 | 1 | 4 | 3
Technical set-up preparation | 0 | 0 | 0 | 1 | 4 | 3
Beamline training | 0 | 0 | 0 | 1 | 4 | 3
Experiment run | 0 | 0 | 0 | 2 | 4 | 2
2013 Statistics kindly provided by User Office

## Slide 33: User feedback
33
Chart: 
| N/A | Unsatisfactory | Poor | Satisfactory | Good | Excellent
Access and support of present laboratories | 0 | 0 | 0 | 1 | 4 | 3
Chart: How does this beamline compare to others elsewhere in the world?
| No basis for comparison | Unfavorably | Favorably
CIRCE | 4 | 0 | 4
2013 Statistics kindly provided by User Office
- “do not allow to do absorption measurements on C K-edge”
- “The beamline status needs to be improved and optimized.”
- I have to emphasize the excellent support from the Circe beamline scientist and beamline staff
- The support of the beam team was excellent.

## Slide 34: Publications and Confs.
Publications in 2013 & 2014:
  One paper about NAPP endstation (J. Phy. Conf. Ser.): (1)
  Other publications by BL staff (still without ALBA data): (14)
2014 Conference contributions and Talks with ALBA data: (13)
Current paper drafts with ALBA data in progress/submitted: (4)
Workshops are organized on annual basis
34

## Slide 35: 1000th ALBA user
Reaching maturity as an infrastructure, indicated by small „coincidences“:
  damaged ICs (discharge) were exchanged in 2 hours
  Fully functional but slow software was discarded for new one
  User comments: „Your crane drivers are more respectful to the experiments than in other synchrotrons“
  User support over telephone
Keep in mind: Continued, strong efforts are needed for further improvements.
From: OkDork by Noah Gordon
35

## Slide 37: Detailed publication list 2013/2014 – CIRCE BL Staff
(With and without ALBA affiliation)
S. Finzio, M. Foerster, et al., Phys. Rev. App. 1, 021001 (2014)
M. Foerster, et al., Appl. Phys. Lett. 104, 072410 (2014)
T. O. Mentes, A. Locatelli, L. Aballe, L.; et al. Ultramicroscopy 130, 82 (2013)
V. Pérez-Dieste, L. Aballe, S. Ferrer, J. Nicolàs, C. Escudero, A. Milán, E. Pellegrin , J. Phys.: Conf. Series 425, 072023 (2013)
S. Finizio, M. Foerster et al., SPIN 03, 1340008 (2013)
C. Escudero et al., Journal of Synchrotron Radiation , 20, 504 -- 508 (2013)
Z. El-Hachemi, C. Escudero et al, Journal of Materials Chemistry C 1, 3337 -- 3346 (2013)
S. Carenco, A. Tuxen, M. Chintapalli, E. Pach, C. Escudero, The Journal of Physical Chemistry C 117, 6259 -- 6266 (2013) .
J. M. Rebled, M. Foerster, et al., Phys. Chem. Chem. Phys., 2013,15, 18274-18280
C. Escudero, M. Salmeron, Surface Science 607, 2 -- 9 (2013)
T. Miyawaki, M. Foerster, et al., J. Appl. Phys. 114, 073905 (2013)
R.M. Reeve, C. Mix, M. König, M. Foerster, G. Jakob, and M. Kläui, Appl. Phys. Lett. 102, 122407 (2013)
Wohlhüter, J. Rhensius, C.A.F. Vaz, J. Heidler, H.S. Körner, A. Bisig, M. Foerster, L. Mechin, F. Gaucher, A. Locatelli, M.A. Nino, A. El Moussaoui, E. Goering, L.J. Heyderman, and M. Kläui, J. Phys.: Condens. Matter 25, 176004 (2013)
A. Tuxen, S. Carenco, M. Chintapalli, C.-H. Chuang, C. Escudero, et al., Journal of the American Chemical Society 135, 2273 -- 2278 (2013)
(Conferences Contributions 2014)
ORAL PRESENTATION Tailor 2014 workshop: N.J. Divins, C. Escudero, V. Pérez-­-Dieste, J. Llorca.
INVITED PRESENATION ICCC-41: J. Veciana, I. Ratera, Souto, M., Morales, D., Pérez-Dieste, V., Escudero, C., et al. .
Abstract IPS-20, V. A. de la Peña O'Shea , L. Collado, J. Fermoso, V. Perez-Dieste, C. Escudero, J. M. Coronado, D. P. Serrano.
6 abstracts for LEEM/PEEM 2014 with ALBA data (L. Martin, T. Schmidt, J. Figuera, L. Aballe, M. Foerster, CIRCE)
Abstract TMRC 2014 (J. Figuera)
Abstract GEFES 2014 (L. Martin)
Talk at FyT 2014 (L. Aballe, CIRCE)
Invited talk at Nanoselect-Consolider meeting 2014 (L. Aballe, CIRCE)