## Slide 1: [Mysterious] vertical beam size measurements using iXD at ALBA
Andriy Nosych
Ubaldo Iriso
DEELS Workshop 2015

## Slide 2: The ALBA machine provides:

## Slide 3: The idea behind IXD:
To provide an alternative-to-pinhole vertical beam size measurement,
possibly in real time
1.4 T
1E4 flux reduction

## Slide 4: Crotch absorbers
Cross-section
bottom jaw
Tooth width: 6 mm
Teeth inclination angle: 8.8o
Ray tracing shows that radiation from previous dipole does NOT pass by the IXD, so we are looking at a single photon source
35 mm of Cu
2 mm of steel
1700 mm
to source

## Slide 5: How it looks like
Storage ring sectors: 16
Dipoles per sector: 2
Absorbers per dipole: 4
Given the space constraints, there are 2 possible IXD locations at the end of each sector. We choose the one closer to the source (second dipole).

## Slide 6: Green-paper study

## Slide 7: Scintillators
Chemical composition: Lu1.8Y0.2SiO5:Ce
Light yield: 32 photons/keV
Decay time: 41 ns
Scintillation emission wavelength: 420 nm
Available size: 1” x 0.8 mm, 1 mm, 2 mm
CRY19 a silicate single crystal for X-ray detectors.
Chemical composition: unknown
Light yield: 28 @300K [103Ph/MeV]
Decay time: 41 ns
Scintillation emission wavelength: 420 nm
Available size: 1” x 1 mm

## Slide 8: Scintillators
Photon absorption by common scintillators (0.8 mm thick)
Simulation by XOP

## Slide 9: Theory & expectation
Vert e-beam size at source (inside dipole): 24 um ()
Vert e-beam size seen by pinhole: 23 um ()
Vert e-beam divergence at source (LOCO): 1.6 urad ()
Photon divergence (XOP simulation) at IXD location: 26 urad ()
Expected vertical photon beam size at IXD:
= 48 um

## Slide 10: Best IXD shots (@10s)
Typical pinhole image
sy =
25 um
sx = 60 um
IXD images
What causes the segmented pattern??

## Slide 11: The teeth imprint
40 um
39 um
6 mm

## Slide 12: Current status (@1s)
FOV: around 18 x 13 mm
Cry19
1 mm
PreLude 2 mm
PreLude 1 mm
sy = 56 um
sy = 57 um
sy = 71 um

## Slide 13: Current status
Measured photon beam size is quite far off the expected value.
Photon divergence corresponding to the measured values is: 32-47 urad
Copper thickness corresponding to this divergence: <5 mm !!!
photon beamsize

## Slide 14: Coupling scans
IXD1 (Sector1) PL 0.8 mm
1.3%
coupling
0.3%
coupling

## Slide 15: Coupling scans
0.3%
coupling
2.4%
coupling

## Slide 16: Screen materials test
 | Lin | PSF* | R2
PL 2mm (1s) | 1 | 25 | .984
PL 2mm (5s) | 1 | 27 | .992
CR 1mm (1s) | 0.8 | 32 | .934
PL .8mm (1s) | 1.2 | 17 | .996
What is enlarging our beam?
PreLude (2 mm)
CRY19 (1mm)
PreLude (1 mm)
PreLude (0.8 mm)
cleanest measurements
brightest
measurements

## Slide 17: Penelope & Fluka simulations
Fluka:
Secondary Cu photons (K_alpha) tracking ON produced 0.2% more beamsize w.r.t. to tracking OFF
Penelope:
Secondary electrons tracking ON
produced 0.1% more beamsize w.r.t to tracking OFF
Neither are able to predict the PSF enlargement of IXD
to evaluate the PSF enlargement by secondary emission

## Slide 18: Vertical bumps
Bumps of 50-100-200-300 um are detected well by the IXD
(calculating the difference in Gaussian centroids)

## Slide 19: Conclusions
We see a linear correlation between beam size measured by IXD and Pinhole
There is a linear offset which is not understood:
- Optics effect? (DOF checked, pixel size calibrated)
- Secondary emission? (checked by tracking simulations: Fluka, Penelope, which
explain the negligible enlargement by secondary photons and electrons)
- Other?
IXD measurements with different screen materials: results are similar but measure different beam sizes, intensities, and have different sensitivity to high energy x-rays.
IXD measurements in different but geometrically-identical machine sectors show difference in photon intensity and IXD image shape.
Difficult to get to sub-second measurement due to overall low residual flux (absorbers are doing a good job).
Difference in attenuation by identical absorber is not fully understood as well

## Slide 20: Thank you!

## Slide 21: Motorized setup (fail)

## Slide 22: More absorption

## Slide 23: Coupling scans

## Slide 24: Backup: IXD calibration
FOV measurement
Alignment in-situ
Pixel size calibration
DOF test (typically 1 mm) so focusing on either side of the screen is OK
Calibration can be cross-checked online by looking at the shadow of marks (if visible, especially at high CCD exposures)

## Slide 25: Abstract
An in-air x-ray detector (IXD) is used for vertical beam size measurements at ALBA. The IXD intercepts the residual x-rays (above 120 keV) which are left of the main flux after passing through copper absorbers. These x-rays generate a visible footprint on a sensitive scintillator and are observed by a CCD camera equipped with simple optics with the exposures in the order of 1-10 seconds. This measurement allows evaluation of the original electron beam size; however, at the moment, several unexpected systematic effects are not understood and forbid the reliable application of IXD for ALBA diagnostics. Among these effects are a) enlarged x-ray fan size w.r.t. theoretical value, b) varying (very low) intensity of the residual x-rays if sampled at different but geometrically identical locations of the machine, c) attenuation effect of the absorber "teeth" on the x-ray fan.

## Slide 26
+Idea + tunnel location + Green paper study
+IXD setup photos
+Scintillators
+Theory, flux, transmission
+Calibration+focusing+line thickness test
+Beam imprints – best picks
+Current status
+ We see different imprints in 3 locations (wtf?)
+ Coupling scans + mismatch with pinhole
+ Coupling scans with changing screens
+ Fluka+XOP+penelope study
+ Bumps
not understood:
enlarged x-ray fan size w.r.t. theoretical value,
varying (very low) intensity at different locations of the SR attenuation effect of the absorber "teeth" on the x-ray fan.