## Slide 1: Solaris Status
Adriana Wawrzyniak
On behalf of Machine Team
National Synchrotron Radiation Center SOLARIS
XXVII ESLS 28.11.19, ALBA-CELLS, Barcelona, Spain

## Slide 2
OUTLINE
Overview
Operation statistics in 2019
Main Failures
New Installations
Projects
Future plans
Summary

## Slide 3
SOLARIS OVERVIEW
PEEM/XAS
U-ARPES
600 MeV Linac with RF Thermionic Gun
6 accelerating structures combined in 3 units
Accelerating gradient 20 MeV/m
S-band – 2998.5 MHz
3 RF Units & SLED cavities
In operation since Dec. 2014
1.5GeV Storage ring
12 DBA Cells – 96 m circ.
Space for ID’s (10 sections) ~ 3.5 m
10 straight sections for IDs
100 MHz RF system
300 MHz Landau Cavities
Injection dipole kicker
Ramping
In operation since May 2015
PHELIX
XMCD
3/26
PINHOLE
LUMOS

## Slide 4
Optics design by S.C. Leemann - MAXIV
MAXIV Facility, DDR, § 3, http://www.maxlab.lu.se/maxlab/max4/DDR_public
Storage Ring Lattice
STORAGE RING PARAMETERS
Parameter | Designed | Measured
Energy | 1.5 GeV | 1.45 ±0.05 GeV
Max. Current | 500 mA | 320 mA (op)
Harmonic number | 32 | 32
Natural emittance (bare lattice) | 5.982 nmrad | 7.5 ±1.5 nmrad
Coupling | 1 % | 0.83 %
Tune νx, νy | 11.22, 3.15 | 11.22, 3.15
Corrected chromaticity ξx, ξy | +2,+2 ; +1, +1 | +1.4, +1.6;+0.9,+0.9
Energy loss/turn | 114.1 keV | 103.7 ±12.3 keV
Momentum acceptance | 4% | 3.7± (0.3)%
Synchronous phase | 168o | 167.4o ± 2.7o
Synchrotron tune | 0.00239 | 0.00228
Physical acceptance h/v | 18 /4 mmrad | 15.68/3.77 mmrad
Lifetime | 13h | 23 h

## Slide 5: Operation statistics in 2019

## Slide 6: 2008 h for Beamlines and 522 h for machine studies. In total 2530h/year of machine operation
OPERATION IN 2019
2 Shifts from Tue to Fri
2 operators on shift
From October shifts 7:00-15:00; 14:00-22:00
2 Injection per day 7:00; 16:00;
On call duties and extra hours during night shifts and Saturdays
Time distribution between machine and beamlines studies
Chart: 
| JANUARY | FEBRUARY | MARCH | APRIL | MAY | JUNE | JULY | AUGUST | SEPTEMBER | OCTOBER | NOVEMBER | DECEMBER
HOURS for BEAMLINES | 0 | 232 | 136 | 256 | 112 | 240 | 192 | 0 | 168 | 256 | 280 | 136
HOURS for MACHINE | 86 | 34 | 56 | 16 | 56 | 32 | 24 | 40 | 82 | 40 | 40 | 16

## Slide 7: Availability=
AVAILABLILITY
 | Availability | MTBF [h] | MTTR[h]
2018 | 90.4% | 16.3 | 1.5
2019 | 91.9% | 23.6 | 1.7
Reliability of UPS with maintenance periods : 99,39%
Reliability of UPS without maintenance periods : 99,42%
Reliability of electrical system with maintenance periods: 97,95%
Reliability of electrical system without maintenance periods: 99,18%

## Slide 8
AVERAGE CURRENT DELIVERED

## Slide 9
MAIN FAILURES
Chart: 
| B | CT | DIA | MPS | ID | INF | I | MML | HE | OI | OT | NET | MAG | PSS | RF | VAC | WAT
 | 0.0065057 | 0.23139 | 0.38773 | 0.0032528 | 0 | 0.082945 | 0 | 0.00054214 | 0.019517 | 0 | 0.17292 | 0 | 0.030143 | 0.0048793 | 0.03795 | 0.018975 | 0.0032528
Chart: 
| B | CT | DIA | MPS | ID | INF | I | MML | HE | OI | OT | NET | MAG | PSS | RF | VAC | WAT
 | 1 | 11 | 30 | 1 | 0 | 2 | 0 | 1 | 3 | 0 | 7 | 0 | 5 | 2 | 15 | 3 | 1
Failures by number
Failures by time

## Slide 10: Main Failures

## Slide 11: Main failures
Problems with injection and accumulation after Spring shutdown (OT)
Problems with orbit corrections and Liberas system (DIA, MML)
Flood in the 1.34 PS room (INF)
Beam losts due to instabilities and plungers (RF)
Problems with SQFI and Semiinstruments PS(MAG)
Problems with correctors magnets(MAG)
Short circuits of PFS (MAG)
100 kVA UPS Failures (INF)
Timing issuses – configurtion after switch off the event generator (CTL)
Communication problems device servers and PS(MAG)
MAIN FAILURES

## Slide 12
19/Nov/2019:
Hard reset of Liberas devices:
- caused loss of one of Libera configuration.
These Liberas which were supposed to control section 1,2 and 3 in ring, were actually controlling sections 6,7 and 8: - reflected in ‘mixed’ response matrix.
20/Nov/2019:
Another consequence of Liberas hard reset:
- caused a confirmation of Liberas hostnames changes introduced 2 weeks earlier.
Once synchronization between astor device servers, Liberas hostnames and IP addresses was done, response matrix has been measured and showed expected pattern.
21/Nov/2019:
Issue with temporary loss of communication between matlab and beam position readings:
- turned out that Java buffer used in Jive was too low: 128MB. Once changed to 4GB, communication issue causing beam drifting was resolved
Main issues with Liberas based on Ubuntu 14.04

## Slide 13: Beam instabilities
Peridical uncorrelated (duration-20min) distorsion of the beam above the accepted limit.
CIRCULATOR FAILURE
Circulator reflected power issue
Beam dumps caused by sudden loss of RF power on Transmitter 2
No signs of reflected power peaks in tango
Reflected power alerts in Tx2 GUI
Extensive testing carried out
Damaged teflon separator replaced with a new one, made to order
New piping installed and soldered in place
Impedance matching with ferrite beads carried out
Final measurements taken every week for about 2 months after repair concluded, that circulator is entirely operational and fixed

## Slide 14
Last failure: 18.09.2019
The SemiInstruments Power Supply went down during the user operation what caused a beam dump. The 3 phases fuse were switched off and PS interlocked. After switching on again it blowed up and went down again. Decision was taken to exchange it to the DELTA ELKTRONIKA PS „SM 15-400”.
PS replacement took about 5 hours.
SQFo_Loc
Diagnostic of failure
The reason for the switch off fuse was a loose (not tightened) 230VAC wire on the contactor as a result of which a spark occurred and the contact has been burned.
Repair
Both contactors and wires has been replaced.
Parts order time: about 1 week
Repair time: about 1 hour
Conclusion
Delta Elektronika PS is less expensive and a stable supply. It still used for SQFo_Loc. Repaired SemiInstruments PS is destinated for a spare.
POWER SUPPLIES FAILURE

## Slide 15: SQFi Last Failure
14.05.2019 - After magnets cycling, SQFi power supply (PS05) had an overcurrent interlock. It was inspected by operator and electrical team, and this may be a sign of device aging and malfunctioning.
17.05.2019 - Interlock on SQFI magnet disabled further operation as it can be reset only via 1.34, which we can't access at this point. Further maintenance planned to next week.
10.06.2019 – Beam lost due to Interlock on SQFi PS „Phase failure” Main Power Switch has been triggered by unknown rason. AC Current ~95[A], Temp. ~35 °C. It was planned to use current analyzer and obsrvation. The new Main Power Switch has been ordered.
11.06.2019 – Main Power Switch on the SQFi Power Supply has been measured. Current on the 3 phases was similar ~100A, but temperature on the 3rd phase was higher than other phases. It was about 100 deg and increasing. The Main Electrician ordered the machine to be turned off. The main switch has been replaced with a substitute. Beam has been lost for about 5 hours.
07.10.2019 – The substitute Main Power Switch has been replaced for the new one without beam lost (machine day). Repair time was about 2 hours. Presumably the reason for the malfunction was the current trigger in the switch, but there are no visible signs of damage. Failure was solved, and ordered one more switch for the spare.
POWER SUPPLIES FAILURE

## Slide 16
UPS 100 KVA FOR IT – CRITICAL DAMAGE 28.06.2019
Report from service :
Following components were damaged:
Tyrystor bypass L1
IGBT booster L1
IGBT inverter L1
Output inverter board L1
Main board power supply
Tape for flex driver IGBT
DC+ Fuse 100A L3
AC Fuse 180A L1
Input Circuit Breaker with auxiliary contact
Logs from UPS: none – damage was too violent.
Probably reasons:
Insects or rodent
Power supply anomaly - high potential on neutral wire or energy return to UPS
Damaged trigger circuit for main circuit breaker DPX 250ER (UPS 100kVA)
Time work without UPS 100kVA for IT- 23 hours
Damage solution:
Borrowed UPS from service ( agreement ) was installed . After 3 weeks new UPS was purchased.

## Slide 17: New Installations

## Slide 18: SOLARIS diagnostic beamlines
PINHOLE DIAGNOSTIC BEAMLINE
The PINHOLE, depicts the electron beam by analyzing the emitted X-rays.
The beamline was installed during summer shutdown 2018.
Now is available to monitor transversal beam size during operation.
LUMOS DIAGNOSTIC BEAMLINE
The LUMOS will be operated in the visible and IR region.
The beamline was installed during summer shutdown 2019.
It is foreseen to take the first beam profile measurements from LUMOS at the end of 2019.

## Slide 19: PINHOLE diagnostic beamline
CONTROL PROGRAM
The PINHOLE software consists
of three layers:
- Taurus based GUI
- TANGO Controls devices
- PLC program
Horizontal (green dots) and vertical (blue dots) emittance measured continuously over 13 hours.
X-ray beam is converted to the visible light at the image plane by 0.2 mm thin scintillator crystal LuAG(Ce) with a peak emission of 535 nm.
Instead of an aluminium exit window a CVD (chemical vapour deposited) diamond with thickness of 400 um was chosen. It improves the simulated photon flux by a factor of 200.
The PINHOLE, depicts the electron beam by analyzing the emitted X-rays.
The beamline was installed during summer shutdown 2018.
Now is available to monitor transversal beam size during operation.
The design of the beamline was modified to provide sufficient X-ray photon flux for proper imaging (CVD diamond window).

## Slide 20: LUMOS diagnostic beamline
The LUMOS will be operated in the visible and IR region.
The beamline was installed during summer shutdown 2019.
It is foreseen to take the first beam profile measurements from LUMOS
at the end of 2019 (next week?).
This beamline will be used for transverse beam profile and longitudinal bunch length measurements as well as to study the longitudinal beam dynamics.
Streak camera model SC-10 with S25 photocathode by Optronis
LUMOS has a similar configuration to the existing diagnostic beamline at MAX IV 1.5 GeV storage ring.
New beamline will enable a wide range of new optical diagnostic measurements.
Adriana Wawrzyniak, Solaris Status, XXVII ESLS Workshop, 26-28.11.19, ALBA-CELLS, Barcelona, Spain

## Slide 21: PHELIX FE commissioning
m/q = 69
beamline open, average pressure at FE: 5.8e-9mbar and BL: 5.4e-9mbar
for a short period of time of conditioning, small contamination was detected at safety shutter in FE area.

## Slide 22: XMCD UVC installation
XMCD
PHELIX
UARPES
STRAIGHT + RGA
STRAIGHT + KICKER
STRAIGHT R1-09
Dipole VK1 + VK2
STRAIGHT R1-08
STRAIGHT ALL
XMCD NEG coated Undulator Vacuum Chamber has been installed by Solaris, after gaining experience from installation of UARPES and PHELIX NEG UVC’s. The lowest value of product P/Ie ~ 1.8e-12 [mbar/mA] calculated based on sputter ion pumps. In case of UARPES next NEG activation is foreseen.
XMCD VACUUM CHAMBER
INSTALLATION

## Slide 23: XMCD FE Installation Undulator Installation
SAT: Test Stand area. Installation and bakeout - two weeks. Then installation and pumping in the storage ring – one week
XMCD FRONT END & UNDULATOR INSTALLATION

## Slide 24: XMCD FE Degassing up to Heat Absorber
Beam current: 362 mA
Energy: 1.5 GeV
Undulator gap: 19 mm
Pressure at Heat Absorber: 1.2e-7 mbar
Temperature at Heat Absorber: ~30 deg. of C
Energy
Beam current
SIP HA
XMCD FE Degassing

## Slide 25: VAC Projects: Construction of Solaris pumping unit (2018/2019)
Due to limited access to the front-end (missing doors) and to avoid long bellows during pumping customize pump station is needed
Three units, so far:
SOLAR 1/2/3
PROJECTS

## Slide 26: VAC Projects: Sophisticated bake-out Heating Unit Controller for synchrotron application (2017)
Features:
Activation of NEG strips in dipole vacuum chambers
Access to machine signals/interlocks during NEG strip activation
Activation of NEG coated undulator vacuum chambers
Activation of NEG strips and NEG coated undulator vacuum chambers in one process
Bake-out of typical vacuum chambers
Protection against burned heating tape
Monitoring of temperature distribution over bake-out system
Cooperation with mobile pump station, pressure interlocks
One button machine
Parameters:
PLC & Tango control system
6 output heating channels (2KW)
2 DC power supplies
PC & Switch
CentOS operating system
Two units, so far:
HUC1 - prototype
HUC2 - copy
PROJECTS

## Slide 27
Tune measurement device
Built according to MAXLAB specification
Stripline feeding network and reciever network operational
Succesful measurements
Analog reflected power interlock for Klystrons
Unreliable performance of Rhode&Shwarz scopes.
False reflected power interlocks – unable to inject beam.
Hardware fix – a dedicated module to detect and signal reflected power levels.
PROJECTS

## Slide 28
UARPES
PEEM/XAS
PHELIX/NAPXPS
XMCD/STXM
SOLAIR
2018
2019
2019
Cryo
EM
MX-SAXS-PD
SOLABS
POLYX
PLANS FOR NEXT BEAMLINES
Operational
Under construction
Operation September 2019 –Users 2020
Pinhole diagnostic beamline operational since september 2018
LUMOS diagnostic beamline to be installed in mid 2019
Construction &
Commissioning

## Slide 29
Daily operation (Injections 7:00, 16:00 current between 280-320mA)
Maintenance activities during shutdowns (active plan for repair – fixing magnets connection, valves connection replacement, short circuit in PFS removal etc.)
Focus on beam stability
Closed orbit correction improvement and Libera firmware and device server upgrade
Working on the python application for the orbit correction
Landau cavities tuning and instabilities supressing (ongoing)
Suppressing impact of the UARPES, PHELIX & XMCD EPUs on the beam position (ongoing)
Installation of the FOFB
Implementation of the tune feedback
Installation and commissioning of the PINHOLE diagnostic beamline done
The installation of the LUMOS diagnostic beamline is almost done, commissioning expected in the beginning of next year
Equipment development (tune measurement device, LLRF test stand, heating unit controller, vacuum pump station)
Installation of the SOLABS FE
Redesign of the VK1 vacuum chamber for FTIR beamline
From 2020 more shifts sheduled for users (40%) and machine development (50%).
SUMMARY & PLANS

## Slide 30: Thank you !
National Synchrotron Radiation Center SOLARIS
Kraków, 10.06.19